## Calibration menu

### Password protection

#### Password insertion screen

The calibration menu is password protected. Prior to entering, the user must enter the password.

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  | C | a | l | i | b | r | a | t | i | o | n |  | m | e | n | u |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  | I | n | s | e | r | t |  | P | I | N |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  | - | - | - | - |  |  |  |  |  |  |  |  |

| 0-9 | Insert number |
| --- | --- |
| ← | Backspace – delete last entered digit |
| ENT | Confirm inserted PIN |
| X | Return to main menu |
| others | No function |

#### Password accepted

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  | C | a | l | i | b | r | a | t | i | o | n |  | m | e | n | u |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  | P | I | N |  | O | K |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| others | No function |
| --- | --- |

No action required. The screen will be left after 3 seconds and the calibration main menu will be entered.

#### Password denied

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  | C | a | l | i | b | r | a | t | i | o | n |  | m | e | n | u |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  | P | I | N |  | D | E | N | I | E | D |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| others | No function |
| --- | --- |

No action required. The screen will be left after 3 seconds and the password insertion screen will appear again.

### Calibration main menu

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | C | a | l | i | b | r | a | t | i | o | n |  | m | e | n | u |  |  |  |  |
| 2 | > |  | 1 | . | S | e | t | p | o | i | n | t |  | c | h | e | c | k |  |  |
| 3 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| ↑ | Move cursor up one line (currently not implemented) |
| --- | --- |
| ↓ | Move cursor down one line (currently not implemented) |
| ENT | Select currently current item |
| X | Leave current menu level and return to main menu |
| others | No function |

The only calibration option now is the verification of setpoint voltages

#### Setpoint check menu

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | S | e | t | p | o | i | n | t |  | c | h | e | c | k |  |  |  |  |  |  |
| 2 | > |  | 1 | . | L | o | w |  | s | e | t | p | o | i | n | t |  |  |  |  |
| 3 |  |  | 2 | . | H | i | g | h |  | s | e | t | p | o | i | n | t |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| ↑ | Move cursor up one line |
| --- | --- |
| ↓ | Move cursor down one line |
| ENT | Select currently current item |
| X | Leave current menu level and return to calibration menu |
| others | No function |

##### Low setpoint check screen

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | L | o | w |  | s | e | t | p | o | i | n | t |  |  |  |  |  |  |  |  |
| 2 | R | o | w | 1 | : |  |  | 6 | 0 | 0 | 0 | V | ¦ |  |  | 1 | 2 | 3 | 4 | V |
| 3 | R | o | w | 2 | : |  |  | 6 | 0 | 0 | 0 | V | ¦ |  |  | 1 | 2 | 3 | 4 | V |
| 4 | R | o | w | 3 | : |  |  | 6 | 0 | 0 | 0 | V | ¦ |  |  | 1 | 2 | 3 | 4 | V |

Where the left value is the low setpoint and the right value is the directly read-back value from the ADC for comparison.

The low setpoint voltage is applied to the outputs automatically as soon as this menu is entered, and a voltage of 0 V is applied to the outputs automatically when this menu is left again.

| X | Leave calibration procedure and return to setpoint check menu |
| --- | --- |
| others | No function |

##### High setpoint check screen

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | H | i | g | h |  | s | e | t | p | o | i | n | t |  |  |  |  |  |  |  |
| 2 | R | o | w | 1 | : |  | 1 | 0 | 0 | 0 | 0 | V | ¦ |  | 1 | 2 | 3 | 4 | 5 | V |
| 3 | R | o | w | 2 | : |  | 1 | 0 | 0 | 0 | 0 | V | ¦ |  | 1 | 2 | 3 | 4 | 5 | V |
| 4 | R | o | w | 3 | : |  | 1 | 0 | 0 | 0 | 0 | V | ¦ |  | 1 | 2 | 3 | 4 | 5 | V |

Where the left value is the high setpoint and the right value is the directly read-back value from the ADC for comparison.

The high setpoint voltage is applied to the outputs automatically as soon as this menu is entered, and a voltage of 0 V is applied to the outputs automatically when this menu is left again.

| X | Leave calibration procedure and return to setpoint check menu |
| --- | --- |
| others | No function |