Photonis 0387 HSB MRTS Tesla repair Software User Manual. Description: User manual for HSB module repair test set; Doc No.:; PHO-0387-S2-2-0. Revision:; 003. Date:; 22-11-2017.

|  | Name | Date | Signature |
| --- | --- | --- | --- |
| Author | M. Mitscherlich | 22-11-2017 |  |
| Reviewed |  |  |  |
| Approved |  |  |  |

* The approval box may be extended as necessary.

No part of this document may be reproduced in any form or by any means, without the prior written permission of the respective owners. All trademarks are property of their respective owners.

Document History

| Version | Date | Status | Description |
| --- | --- | --- | --- |
| 001 | 01-09-2017 | draft | Initial draft |
| 002 | 07-09-2017 | draft | Updates after review by KRe and GRo |
| 003 | 22-11-2017 | draft | Updates after review by Photonis Updates after spec additions |

Table of Contents

1	Introduction	4

1.1	Scope	4

1.2	References	4

1.3	Definitions	4

2	System overview	5

2.1	Display	5

2.2	Keypad	6

3	Service manual	7

3.1	Selection cursor	7

3.2	Start screen	8

3.3	Main menu	9

3.3.1	Tesla repair - Tube repair menu	10

3.3.1.1	Select preset menu	10

3.3.1.2	Preset preview screen	10

3.3.1.3	Start repair menu	11

3.3.1.3.1	Error messages based on safety	11

3.3.1.3.2	Repair in progress	11

3.3.1.4	Repair pause	12

3.3.1.4.1	Repair paused	12

3.4	Administrator Menu	13

3.4.1	Password protection	13

3.4.1.1	Password insertion screen	13

3.4.1.2	Password accepted	13

3.4.1.3	Password denied	13

3.4.2	Operator main menu	14

3.4.2.1	Change PIN menu	15

3.4.2.1.1	Insert new Pin	15

3.4.2.1.2	New Pin OK	15

3.4.2.1.3	New Pin failed	16

3.4.2.2	I/O control menu	17

3.4.2.2.1	Read interlock	18

3.4.2.2.2	Solenoids	19

3.4.2.2.3	Get voltage in	20

3.4.2.2.4	Teslagun relais	22

3.4.2.3	Information and version screen	23

3.5	Calibration menu	24

3.5.1	Password protection	24

3.5.1.1	Password insertion screen	24

3.5.1.2	Password accepted	24

3.5.1.3	Password denied	24

3.5.2	Calibration main menu	25

3.5.2.1	Calibrate ADCs menu	25

3.5.2.1.1	Calibration procedure for an ADC channel	26

# Introduction

## Scope

This document describes the end-user functionality of the Photonis HSB module repair test set. It gives detailed overview over the menu structure and explains the interactivity between user and device.The initial status of this document will be Preliminary, when it acts as an internal and external discussion document. The first formal and complete document, which will be send to the customer, will be marked as Draft, and will have revision level 1.0. This document will serve as formal document for review purposes. If both Micro-key and the customer approve of the contents of the document, it will be marked as Final.

## References

Document somename, vvv
Micro-Key, dd-mm-yyyy

## Definitions

TBD		To Be Defined

# System overview

Interaction between user and system occurs via two devices which are part of the HSB module repair test set.

## Display

Visual feedback and information is provided to the user through a display. It contains four rows with 20 characters each.

Figure 1: New heaven display NHD-0420D3Z-FL-GBW-V3

Within this document the display content is related to by a table. Below represents the display seen from left to right and up to down as a matrix of 4x20.

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

As example, the table below represents the same display content as in Figure 1.

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | S | c | h | e | d | u | l | e | : |  | M | o | n | . | - |  | F | r | i | . |
| 2 | O | p | t | i | o | n | s |  | P | r | e | s | e | t | : |  | A | 2 | 1 |  |
| 3 | A | u | t | o | - | o | f | f | : |  | N | O |  |  |  |  |  |  |  |  |
| 4 | A | c | c | e | s | s | : |  | P | u | b | l | i | c |  |  |  |  |  |  |

## Keypad

The system contains a keypad which represents the only possibility for the user to interact with the several functions of the system.

Figure 2: 16-key keypad

The keypad consists of a matrix of four by four keys, making it 16 keys totally. They are labelled as follows:

| 0-9 | Numeric keys 0 - 9 |
| --- | --- |
| ← | Arrow left |
| ↑ | Arrow up |
| → | Arrow right |
| ↓ | Arrow down |
| ENT | Enter |
| X | Escape |

# Service manual

## Selection cursor

As soon as a menu with selectable items appears on the display there must be a cursor/selector to differ between the currently highlighted menu item and other items.

The cursor is character: >

The cursor is always in first position of the current row, followed by a blank (space character).

For example, the cursor active in third row with random text behind looks like:

| / | 1 | 2 | 3 | 4 |
| --- | --- | --- | --- | --- |
| 1 | R | a | n | d |
| 2 |  |  | 1 | . |
| 3 | > |  | 2 | . |
| 4 |  |  | 3 | . |