## Calibration menu

### Password protection

#### Password insertion screen

The calibration menu is password protected. Prior to entering, the user must enter the password.

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  | C | a | l | i | b | r | a | t | i | o | n |  | m | e | n | u |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  | I | n | s | e | r | t |  | P | I | N |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  | - | - | - | - |  |  |  |  |  |  |  |  |

| 0-9 | Insert number |
| --- | --- |
| ← | Backspace – delete last entered digit |
| ENT | Confirm inserted PIN |
| X | Return to main menu |
| others | No function |

#### Password accepted

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  | C | a | l | i | b | r | a | t | i | o | n |  | m | e | n | u |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  | P | I | N |  | O | K |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| others | No function |
| --- | --- |

No action required. The screen will be left after 3 seconds and the calibration main menu will be entered.

#### Password denied

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  | C | a | l | i | b | r | a | t | i | o | n |  | m | e | n | u |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  | P | I | N |  | D | E | N | I | E | D |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| others | No function |
| --- | --- |

No action required. The screen will be left after 3 seconds and the password insertion screen will appear again.

### Calibration main menu

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | C | a | l | i | b | r | a | t | i | o | n |  | m | e | n | u |  |  |  |  |
| 2 | > |  | 1 | . |  | C | a | l | i | b | r | a | t | e |  | A | D | C | s |  |
| 3 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| ↑ | Move cursor up one line (currently not implemented) |
| --- | --- |
| ↓ | Move cursor down one line (currently not implemented) |
| ENT | Select currently current item |
| X | Leave current menu level and return to main menu |
| others | No function |

The only calibration option at the moment is the analogue-to-digital converter, or analogue input.

#### Calibrate ADCs menu

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | C | a | l | i | b | r | a | t | e |  | A | D | C | s |  |  |  |  |  |  |
| 2 | > |  | 1 | . |  | C | h | a | n | n | e | l | 1 |  |  |  |  |  |  |  |
| 3 |  |  | 2 | . |  | C | h | a | n | n | e | l | 2 |  |  |  |  |  |  |  |
| 4 |  |  | 3 | . |  | C | h | a | n | n | e | l | 3 |  |  |  |  |  |  |  |

| ↑ | Move cursor up one line |
| --- | --- |
| ↓ | Move cursor down one line |
| ENT | Select currently current item |
| X | Leave current menu level and return to main menu |
| others | No function |

##### Calibration procedure for an ADC channel

The following menu is similar for all ADC channels

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | C | a | l | i | b | r | a | t | e |  | A | D | C |  | # | 1 |  |  |  |  |
| 2 |  |  |  |  |  |  | P | r | o | v | i | d | e |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  | 0 |  | k | V |  |  |  |  |  |  |  |
| 4 |  |  |  |  | a | n | d |  | h | i | t |  | E | N | T |  |  |  |  |  |

| ENT | Confirm provided voltage and continue to next calibration point |
| --- | --- |
| X | Leave calibration procedure and return to calibrate ADCs menu |
| others | No function |

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | C | a | l | i | b | r | a | t | e |  | A | D | C |  | # | 1 |  |  |  |  |
| 2 |  |  |  |  |  |  | P | r | o | v | i | d | e |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  | 1 |  | k | V |  |  |  |  |  |  |  |
| 4 |  |  |  |  | a | n | d |  | h | i | t |  | E | N | T |  |  |  |  |  |

| ENT | Confirm provided voltage and continue to next calibration point |
| --- | --- |
| X | Leave calibration procedure and return to calibrate ADCs menu |
| others | No function |

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | C | a | l | i | b | r | a | t | e |  | A | D | C |  | # | 1 |  |  |  |  |
| 2 |  |  |  |  |  |  | P | r | o | v | i | d | e |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  | 1 | 0 |  | k | V |  |  |  |  |  |  |  |
| 4 |  |  |  |  | a | n | d |  | h | i | t |  | E | N | T |  |  |  |  |  |

| ENT | Confirm provided voltage and continue to next calibration point |
| --- | --- |
| X | Leave calibration procedure and return to calibrate ADCs menu |
| others | No function |

| / | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | C | a | l | i | b | r | a | t | e |  | A | D | C |  | # | 1 |  |  |  |  |
| 2 |  |  |  |  | C | a | l | i | b | r | a | t | i | o | n |  |  |  |  |  |
| 3 |  |  |  |  |  |  | f | i | n | i | s | h | e | d |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| ENT | Confirm and return to calibrate ADCs menu |
| --- | --- |
| others | No function |

The exact calibration procedure is not defined, yet, so the menu definition, especially the calibration points mentioned, are subject to changes!